## Slide 1
  G460 PID sensor
  performance and maintenance
  GfG Instrumentation, Inc.
  1194 Oak Valley Drive, Suite 20
  Ann Arbor, Michigan 48108
  Toll free (USA and Canada): (800) 959-0329
  Direct: +1-734-769-0573
  Service e-mail: service@gfg-inc.com
  Internet: www.gfg-inc.com
  Technical documentation and download site: www.goodforgas.com

## Slide 2: G450 / G460 Advanced Service Procedures
WARNING: Advanced service procedures should only be undertaken by authorized personnel

## Slide 3: G460 Smart Sensor PID
PID Smart-Sensor
  Broad range VOC measurement
  Extremely sensitive
  Available in two ranges:
    0.5 – 2,000 ppm (Standard PID sensor)
    0.1 – 500 ppm (Optional high resolution PID sensor)

## Slide 4: PID range and resolution
Two versions of the PID sensor available for G460:
  “Standard” PID provides 0.5 ppm resolution over 0 – 2000 ppm (isobutylene scale)
  “High Resolution” PID provides 0.1 ppm resolution over 0 – 500 ppm (isobutylene scale)
“VOC” choice allows the user to specify custom correction factor for a gas not included in the standard on-board library
The full range for the gas selected depends on the relative response of the sensor to the target gas compared to isobutylene
  For instance, when “NH3” (ammonia) is selected, because of the lower relative response to ammonia compared to isobutylene, the full range is expanded from 0 – 2000 (iso scale) to 0 – 6000 ppm (NH3 scale)

## Slide 5: PID sensor menu
PID sensor choices include "Range and Gas"
Use to choose correction factor for new gas from PID library
PID readings displayed in measurement units of gas selected
Name of gas selected will appear in the sensor menu PID position
In normal operation screen will show name of new gas
Range and Gas
0-2000ppm iButyln
0-2000ppm Gasolin
0-2000ppm MIBK
0-2000ppm Acetone
0-2000ppm Deether
0-2000ppm Propyln
0-2000ppm iButyln
CH4
CO
iButyln

## Slide 6: PID sensor “Gas and Unit” library choices
PID Gas List Abbreviations | Common Name | Range with 0 – 2000 ppm full range PID (ISO) | Range with 0 – 500 ppm full range PID (ISO)
iButyln | Isobutylene | 0 – 2000 | 0 – 500
VOC | Generic VOC with user assigned CF | 0 – 2000 | 0 – 500
Gasolin | Gasoline | 0 – 2000 | 0 – 500
MIBK | Methyl-iso-butyl-ketone | 0 – 2000 | 0 – 500
Acetone | Acetone | 0 – 2000 | 0 – 500
Deether | Diethylether | 0 – 2000 | 0 – 500
Propyln | Propylene | 0 – 2000 | 0 – 500
MEK | Methyl-ethyl-ketone | 0 – 1500 | 0 – 375
Diesel | Diesel | 0 – 1500 | 0 – 375
TrClEyn | Trichloroethylene | 0 – 1000 | 0 – 250
Benzene | Benzene | 0 – 1000 | 0 – 250
Toluene | Toluene | 0 – 1000 | 0 – 250
Xylene | Xylene | 0 – 1000 | 0 – 250
Styrene | Styrene | 0 – 800 | 0 – 200
Jetfuel | Jet fuel (JP-8) | 0 – 800 | 0 – 200
nButnol | n-Butyl-alcohol | 0 – 6000 | 0 – 1500
EtActat | Ethyl acetate | 0 – 6000 | 0 – 1500
nHexane | n-Hexane | 0 – 6000 | 0 – 1500
NH3 | Ammonia | 0 – 6000 | 0 – 1500
cHexane | Cyclo hexane | 0 – 3000 | 0 – 750
VyChlrd | Vinyl chloride (VCM) | 0 – 3000 | 0 – 750
MeBromd | Methyl bromide | 0 – 3000 | 0 – 750
nNonane | n-Nonane | 0 – 3000 | 0 – 750
Octane | Octane | 0 – 3000 | 0 – 750
Heptane | Heptane | 0 - 3000 | 0 – 750

## Slide 7: G460 Interchangeable Smart Sensors
Five Smart Sensor positions on PCB:
All you need to do is plug the sensor into a position designed for that type of sensor
  EC 1: COSH
  EC 1, 2, 3: CO, H2S, O2, NH3, SO2, H2, PH3, HCN
  EC 2, 3: NO, NO2, CL2, HCL, ETO, O3, ClO2, HF
  EC 2: PID
  PL: 1 – 100% LEL “pellistor” sensor
  IR: 0.1 - 5.0 Vol % CO2; 0 – 100% LEL combustible; 0 – 100% vol combustible

## Slide 8: G460 Main PCB: Five Smart Sensor Positions
EC1 (Lower Left)
EC2 / PID (Lower Right)
EC3 (Upper Left)
IR (Center)
PL (Upper Right)

## Slide 9: G460 PID maintenance
The G460 PID is protected by both external and internal filters
  Because gas diffuses into and out of the sensor (rather than using a pump to pull the atmosphere across the lamp and electrodes) the system is less prone to particulate contamination
        Note: The PID lamp and electrodes
        should only be cleaned when needed!
The primary symptoms that indicate the need to clean the lamp are:
    Unstable readings
    Oversensitivity to humidity
    Failure to calibrate

## Slide 10: G460 PID Maintenance
The following slides show the step-by-step procedure for cleaning the G460 PID
The manufacturer of the PID lamp (Baseline-MOCON, Inc.) has also posted a training video at the following link:
    http://vimeopro.com/gasanalysis/pid-tech-plus
    The password is: baseline11
G460 PID cleaning kit (PN 7740-026): includes replacement filters, lamp polishing pads, tweezers and compressor stick
        Note: The appearance of the Baseline-MOCON
        “PID Plus” sensor in the video is slightly different
        from the GfG PID version. The procedure for
        disassembly and cleaning the lamp is exactly the
        same, however.

## Slide 11: G460 PID Maintenance Kit
Lamp polishing pad
Padded stick
Outer filter (moisture barrier)
Inner filter
PID maintenance kit components

## Slide 12: G460 PID Maintenance Cautions
Wear gloves when handling or disassembling PID
    Note: Direct contact between fingers and PID lamp, electrodes and other components can leave oils and contaminants behind that can degrade performance

## Slide 13: G460 PID Maintenance
Make sure instrument is turned off!
Remove battery, open housing, and CAREFULLY remove main board and display assembly exposing sensors
Remove PID sensor
PID sensor

## Slide 14: G460 PID Maintenance
Use tweezers to pry the top off of the sensor
Position the tweezers next to the cap opening
Remove the sensor cap

## Slide 15: G460 PID Maintenance
Remove filters (inner and outer)
Remove spacer

## Slide 16: G460 PID Maintenance
Gently remove sensor PCB
Grip at back of board (near pins)

## Slide 17: G460 PID Maintenance
Use tweezers to GENTLY pry the lamp out of the sensor
Do not touch window or body of lamp with naked fingers

## Slide 18: G460 PID Maintenance
Use circular motion to polish face of lamp window with polishing pad

## Slide 19: G460 PID Maintenance
Clean any lamp, pin, or electrode PCB surfaces that have come into contact with naked skin with alcohol before reassembling
Make sure that all components are COMPLETELY air-dried before reassembly
Do not use blowers or heated air sources to speed up drying!

## Slide 20: G460 PID Maintenance
Metal pads
Metal contacts
Insert lamp back into PID sensor assembly
    NOTE: Metal pads in PID lamp MUST line up with contacts in socket of sensor assembly

## Slide 21: G460 PID Maintenance
Use padded stick to press lamp into place

## Slide 22: G460 PID Maintenance
Reassemble:
  Sensor PCB
  Spacer
  Filters (2)
  Sensor cap
Plug PID sensor back into instrument
    Calibrate the PID sensor before returning the G460 to service!

## Slide 23: Returning the instrument to service
Calibrate ALL sensors in the instrument (whether or not they have been changed) before returning the instrument to service
It is best to let new sensors stabilize in the instrument for 30 minutes prior to calibration

## Slide 24: Questions?