## Slide 1
  G450 / G460
  Updating instrument firmware
  GfG Instrumentation, Inc.
  1194 Oak Valley Drive, Suite 20
  Ann Arbor, Michigan 48108
  Toll free (USA and Canada): (800) 959-0329
  Direct: +1-734-769-0573
  Service e-mail: service@gfg-inc.com
  Internet: www.gfg-inc.com
  Technical documentation and download site: www.goodforgas.com

## Slide 2: G450 / G460 Advanced Service Procedures
WARNING: Advanced service procedures should only be undertaken by authorized personnel

## Slide 3: Updating G450 / G460 Firmware
You will need:
  Latest firmware version to be installed
    Latest version can be downloaded from www.goodforgas.com
  Charger cradle and data download interface cable
    Make sure the GfG drivers for the interface cable have been installed on your computer

## Slide 4: Update firmware: Step 1
Load the new “Update firmware” application onto your computer
  The latest version can be downloaded at www.goodforgas.com
Read the description of changes from the previous version
Decide whether to proceed with update

## Slide 5: Update firmware: Step 2
Connect download interface cable to cradle and connect cable to USB port in computer
Turn instrument on then place in cradle
  Make sure neither the instrument or computer run out of power during update procedure
  Make sure battery gauge in LCD shows at least two segments
  Do not disconnect or disturb interface cable and instrument during update process
Battery power gauge

## Slide 6: Update firmware: Step 3
Open G450 / G460 Update program
Select com port from pull down menu then click on “connect” button
  Note: Correct com port is usually the one with the highest number
Choose desired set of languages from pull-down menu
  G450 supports sets of three languages at a time as choices in on-board memory
  G460 supports all available languages at same time in on-board memory
Com Port Connect Button
Pull-down G450 language menu

## Slide 7: Update firmware: Step 4
“Start Update” Button
Click on “Start Update” button
  A green status bar will indicate progress
The instrument will turn itself back on in normal operation when the update is complete
  The instrument must be calibrated before it is put back into service!

## Slide 8: Update firmware: If a problem is encountered
A red status bar indicates a problem has been encountered
Remove the instrument from the cradle, then replace in cradle and click on the start button to repeat the update process
  If the update application fails to restart the procedure, try momentarily disconnecting the interface cable from the computer, and / or closing and restarting the update application

## Slide 9: Returning the instrument to service
Calibrate ALL sensors in the instrument (whether or not they have been changed) before returning the instrument to service
It is best to let new sensors stabilize in the instrument for 30 minutes prior to calibration

## Slide 10: Questions?