## Slide 1
  GfG Instrumentation, Inc.
  GfG Instrumentation, Inc.
  1194 Oak Valley Drive, Suite 20
  Ann Arbor, Michigan 48108
  Toll free (USA and Canada): (800) 959-0329
  Direct: +1-734-769-0573
  Service e-mail: service@gfg-inc.com
  Internet: www.goodforgas.com

## Slide 2
GfG Instrumentation
World-wide manufacturer of gas detection solutions

## Slide 3: One of the World’s Leading Manufacturers of Gas Detection Products

## Slide 4: GfG Instrumentation mbH
Worldwide headquarters in Dortmund, Germany
An industry leader in development and production of gas measurement technology for over 50 years
Founded in 1961
Over 300 employees worldwide in 2012

## Slide 5: GFG Instrumentation, Inc.
Headquarters in Ann Arbor, Michigan, USA
Responsible for manufacturing, design, and sales support for North and South America
Originally founded in 1963 as Dynamation, Inc.

## Slide 6
GfG Instrumentation
Exceptional designs with best cost of ownership in the gas detection industry
September 12, 2014 GfG Instrumentation company overview
Slide 6

## Slide 7: GfG Product Applications
Confined space
Fire Service
Emergency response
Oil industry
Chemical industry
Steel industry
Water / waste water
Sewer entry
Telecommunications
Municipal departments
Refrigeration systems
Power generation
Pulp and paper

## Slide 8: Portable gas detectors
G460
1 - 6 gas detector
G450
4-gas detector
Micro IV
single-gas detector

## Slide 9: MICRO IV Single Sensor Gas Detector
Compact and affordable single gas detector
Available for wide range of toxic gases, hydrogen or oxygen
Stores long-term and short-term average values (TWA, STEL)
Event logger standard
Data logger standard
Display backlight standard
Vibrator alarm standard
IR Communication Interface

## Slide 10: MICRO IV Single Sensor Gas Detector
Small and lightweight
Extremely robust
Very loud buzzer 95 dB(A)
Bright alarm
Easy and fast sensor replacement
Up to 6 month continuous operation with one battery
Leave turned on, or turn off when not in use to extend battery life
3 alarm thresholds per gas
Certification: c-UL-us Intrinsically Safe for use in Class I, Div. 1, Groups A, B, C, and D Hazardous Locations

## Slide 11: MICRO IV Single Sensor Gas Detector
Display shows in large letters alternating:
  Current gas and gas concentration
  Detection range
Confidence beep adjustable
Keys for Switching ON/OFF
Custom setup with configuration software

## Slide 12: Unique attachable Micro IV motorized pump
Simple slide on design
Add to any Micro IV whenever needed
Pump powered by its own alkaline battery

## Slide 13: Data-logging standard!
Logging of all event data
  Date and time of alarm
  Type of alarm
  Gas concentration which triggered alarm
Consistent storage of data
Infrared interface for data transfer via DI 220 data transfer station

## Slide 14: DI-220 Single Unit Docking / Data Transfer Station
Use for:
  Calibration and bump testing
  Instrument setup and programming
  Download monitoring results to PC

## Slide 15: Automatically records exposure events
No Date Time Type Value STEL(15min) TWA(8h) Gas BAT MK
1 28.08.2006 16:05 Dev-OFF 85% 343-05 48563
2 28.08.2006 16:08 Dev-ON 90% 343-05 48563
3 28.08.2006 16:13 A1 ON 33 0 0 ppm CO 90%
4 28.08.2006 16:13 A2 ON 73 0 0 ppm CO 75%
5 28.08.2006 16:13 A1 OFF 29 1 0 ppm CO 85%
6 28.08.2006 16:14 A2 OFF 0 1 0 ppm CO 80%
No Date Time Type Value STEL(15min) TWA(8h) Gas BAT MK
1 28.08.2006 16:05 Dev-OFF 65% 343-05 48568
2 28.08.2006 16:08 Dev-ON 70% 343-05 48568
3 28.08.2006 16:13 A1 ON 54 0 0 ppm CO 70%
4 28.08.2006 16:13 A1 OFF 25 0 0 ppm CO 55%
Event Data Device 2
Event Data Device 1

## Slide 16: Automatically records time stamped monitoring results
DV-SN DV-ID Sensor BAT Ver Conf. Blip DS-SN DS-Ver
2 GfG MK343-05/48568 70 2.14 OFF 6070887 2.06
4 MK343-05/48563 90 2.14 OFF 6070887 2.06
6 GfG MK343-05/48567 95 2.14 OFF 6070887 2.06
5 GfG MK343-05/48570 45 2.14 OFF 6070887 2.06
5 GfG MK343-05/48570 45 2.14 OFF 6070887 2.06
2 GfG MK343-05/48568 70 2.14 OFF 6070887 2.06
6 GfG MK343-05/48567 95 2.14 OFF 6070887 2.06
4 MK343-05/48563 90 2.14 OFF 6070887 2.06
Test Gas Gas Bottle Test Date Time Mode Test Ready
200 ppm CO IC-39MNRC 28.08.2006 15:52 Bump Yes
200 ppm CO IC-39MNRC 28.08.2006 15:52 Bump Yes
200 ppm CO IC-39MNRC 28.08.2006 15:52 Bump Yes
200 ppm CO IC-39MNRC 28.08.2006 15:52 Bump Yes
200 ppm CO IC-39MNRC 28.08.2006 15:57 CAL Yes
200 ppm CO IC-39MNRC 28.08.2006 15:57 CAL Yes
200 ppm CO IC-39MNRC 28.08.2006 15:57 CAL Yes
200 ppm CO IC-39MNRC 28.08.2006 15:57 CAL Yes
DS220 Data
Test parameter

## Slide 17: Setup choices
Menu
Copyright GfG 2010 17

## Slide 18: Docking Station Micro IV DS-220
Automatic Teast Station for Micro IV
Auto Bump test
Auto Calibration
Bump or calibrate up to six instruments at the same time
Test results automatically results stored on MMC flash memory card

## Slide 19: DS 220 Docking “Station
Complete device management system:
  Automatic bump test
  Automaric calibration adjustment
  Programming interface
  Data transfer
Easy operation
  Push green button for “Bump Test”
  Push red button for “Auto Calibration” adjustment
Test up to 6 detectors simultaneously
No PC required!

## Slide 20: DS 220 Docking Station
Automatically stores results:
  Battery capacity
  Alarm thresholds
  Visual alarm
  Audible alarm
  Serial number (detector)
  Sensor serial number
  Software version
  Detector identification
  Faults
  Alarm list
  “Event record” with gas concentration and alarm status
  Response time for alarm 1 and alarm 2

## Slide 21: Test Results
LED indicators signal test results:
  Green = Detector OK
  Yellow = Test in process
  Red = Test failed

## Slide 22: Calibration adjustment
Press red key to start automatic calibration
Up to 6 instruments can be adjusted simultaneously
No mishandling possible
Shorter waiting time, faster completion!

## Slide 23: Automatically store test and calibration results
Horn Error T-A1 T-A2 T-50% ZPT-before SPAN-before ZPT-akt SPAN-akt
Ok 3,2 4,2 7,2
Ok 2,7 3,7 5,7
Ok 2,4 3,4 6,4
Ok 1,4 3,4 5,4
3,9 1820,4 -0,9 1886,8
1,6 1888,0 -3,1 2050,8
-4,1 1701,6 -7,6 1761,6
-5,6 1713,6 -9,3 1772.0
Test result
BUMP test
Adjustment

## Slide 24: MICRO IV Single Sensor Gas Detector
Available for use with a wide range of dependable, substance specific electrochemical sensors
O2, CO and H2S sensors warranted for 3-full years
Optional range and resolution available for many sensors
Plug-and-play sensor design makes changing or replacing sensors a snap

## Slide 25: Substance-specific electrochemical sensors
Gas diffusing into sensor reacts at surface of the sensing electrode
Sensing electrode made to catalyze a specific reaction
Use of selective external filters further limits cross sensitivity

## Slide 26: Available electrochemical sensors, standard ranges and resolution
More types of EC sensors available every year, both for individual toxic gases as well as sensors designed to detect a range of toxic or combustible gases

## Slide 27: Additional gases detectable by means of relative response
Electrochemical sensors are designed with specific usage requirements in mind
The same manufacturer may offer multiple models of sensor for the detection of the same gas, but that are optimized for different sets of interferents and operating conditions
Thus, cross sensitivities may vary widely between different models and brands of sensors!
In addition, response values may differ at concentrations other than the ones listed in product documentation
Discuss with manufacturer BEFORE attempting to use relative response values to measure additional gases

## Slide 28: Typical Electrochemical Detection Mechanism
H2S Sensor:
Hydrogen sulfide is oxidized at the sensing electrode:
H2S + 4H2O H2 SO4 + 8H+ + 8e-
The counter electrode acts to balance out the reaction at the sensing electrode by reducing oxygen present in the air to water:
2O2 + 8H+ + 8e- 4H2O
And the overall reaction is: H2S + 2O2 H2 SO4
4HS Signal Output: 0.7  A / ppm H2S

## Slide 29: Electrochemical Sensor Performance

## Slide 30: Notable recent gas and vapor TLV® updates
Hydrogen sulfide (2010)
Sulfur dioxide (2009)
Nitrogen dioxide (2012)

## Slide 31: Exposure limits for H2S
4
4
0
Old TLV:
  TWA = 10 ppm
  STEL = 15 ppm
New (2010) TLV:
  TWA = 1.0 ppm
  STEL = 5.0 ppm

## Slide 32: Are H2S sensors capable of measuring at the new TLV limits?
The answer is “Yes” BUT with qualifications…..
  Depends on the manufacturer!
  Sensor must be capable of providing 0.1 or 0.2 ppm resolution
  Instrument design must permit setting the alarms at the desired concentration
  Micro IV designed for measurement at TLV concentration limits!

## Slide 33: Where should you set the H2S alarms?
H2S TLV® only includes STEL and TWA limits; does not include a Ceiling or “Peak” limit
Instruments typically have four user settable alarms for each toxic sensor (Low, High, STEL and TWA)
Suggested alarms:
  NIOSH:
    Low: 10.0 ppm
    High: 15.0 ppm
    STEL: 15.0 ppm
    TWA: 10.0 ppm
  TLV®:
    Low: 3.0 ppm
    High: 5.0 ppm
    STEL: 5.0 ppm
    TWA: 1.0 ppm

## Slide 34: Exposure limits for SO2
OSHA PEL:
  TWA = 5.0 ppm
NIOSH REL:
  TWA = 2.0 ppm
  STEL = 5.0 ppm
Old TLV :
  TWA = 2 ppm
  STEL = 5 ppm
New (2009) TLV:
  STEL = 0.25 ppm

## Slide 35: Suggested alarm settings for SO2
Suggested alarms:
  NIOSH:
    Low: 2.0 ppm
    High: 5.0 ppm
    STEL: 5.0 ppm
    TWA: 2.0 ppm
  TLV®:
    Low: 0.75 ppm
    High: 1.25 ppm
    STEL: 0.25 ppm
    TWA: 0.25 ppm

## Slide 36: Exposure limits for NO2
US OSHA PEL:
  Ceiling = 5 ppm
US NIOSH REL:
  15 min. STEL = 1 ppm
Old TLV:
  8 hr. TWA = 3 ppm
  5 min. STEL = 5 ppm
New 2012 TLV
  8 hr. TWA = 0.2 ppm

## Slide 37: Suggested alarm settings for NO2
Suggested alarms:
  NIOSH:
    Low: 3.0 ppm
    High: 5.0 ppm
    STEL: 1.0 ppm
    TWA: 1.0 ppm
  TLV®:
    Low: 0.6 ppm
    High: 1.0 ppm
    STEL: 0.2 ppm
    TWA: 0.2 ppm

## Slide 38: Sales Support: www.Goodforgas.com
Data sheets
Price lists
Manuals
Application Notes
Product images
Print ads
...and more!

## Slide 39: Questions?